侵入式脑机接口芯片的研发

杨杰，Mohamad Sawan

西湖大学工学院先进神经芯片中心，杭州 310030

人类与数字世界之间的通讯渠道经历了不断的更迭和演进，从打孔卡、键盘、鼠标，到触摸屏和自然语言处理，通讯效率不断提高。这种通讯的本质是人类大脑或神经系统与数字世界的信息交流。到了21世纪的今天，建立大脑和数字世界之间的直接通讯成为了人类的最终目标之一。如今，侵入式脑机接口技术借助先进的集成电路芯片，在建立这种直接通讯方面取得了重要进展。通过植入式电极以及信号记录、数据处理、供能、数据传输与刺激等关键电路模块，侵入式脑机接口为人脑与数字世界之间搭建了双向沟通桥梁。这项技术为脑疾病的诊断和治疗、功能康复等重要医疗应用提供了实现可能[1, 2]。在侵入式脑机接口中，集成电路设计技术扮演着重要的角色。通过将关键电路部分或全部集成在平方毫米级别大小、毫瓦甚至更低功耗的半导体芯片上，摆脱了尺寸、功耗、电缆连接等方面的限制，从而使得侵入式脑机接口系统更容易、更安全、更舒适地植入。

本章将从脑机接口芯片设计的主要考虑出发，介绍脑机接口芯片设计的关键问题和技术挑战，并对主要的电路模块设计进行梳理，以帮助读者全面系统地了解脑机接口芯片的研发。最后，本章将展望脑机接口芯片未来的发展方向。

- 侵入式脑机接口芯片概述

图8-1 侵入式脑机接口芯片总体结构

侵入式脑机接口芯片主要完成脑组织内电极信号的放大、采样、数据处理以及刺激功能，其主要组成如图8-1所示。其中，前置放大器负责放大微弱神经电信号，采样与转换电路对放大后的模拟信号进行采样和数字化，数据处理单元对数字信号进行数据压缩、特征提取以及处理分析等工作，刺激模块利用电刺激方式进行对神经系统的调控。此外，无线数据传输模块能够通过无线方式传输神经信号或特征，避免有线连接带来的创口。无线能量传输模块则通过外部供能降低或消除系统对电池的需求。

如图8-2所示，侵入式脑机接口根据电极形式以及植入位置的不同，通常可采集皮层脑电图（Electrocorticography，ECoG），局部场电位（local field potential，LFP）或者动作电位（action potential，AP）[3, 4]。ECoG信号的电压幅值在0.01至1 mV，频率范围在0.01至200Hz范围。LFP的电压幅值在0.5 mV至5 mV范围，频率范围在0.01Hz至500Hz。AP信号的电压幅值在0.01 mV至5mV，频率范围在100Hz至10k范围[5]。芯片中模拟前端放大器根据系统需求的不同，需要对感兴趣的信号进行放大，并同时抑制相关噪声。

经放大器放大后，神经信号通过模数转换器（analog to digital converter，ADC）转换为更具有抗噪特性的数字信号。由于脑机接口系统一般有数十至上百通道，伴随而来的是大量数据的产生。以100通道，20kHz采样率和10比特采样精度的典型系统为例，系统数据率将达到20Mbps。因此，高速且低功耗的数据传输电路对侵入式脑机接口尤为重要。为缓解带宽压力，可以对信号进行数据压缩，例如直接提取脉冲信号而非保留所有原始信号，或者仅提取数据的特征。

图8-2 侵入式脑机接口信号以及电极特点

侵入式脑机接口芯片的主要优势在于可以在保证数据采集和传输等质量的前提下，降低脑机接口系统的功耗和尺寸，并提高系统可采集的神经通道数量。最近几年发表的脑机接口芯片单个通道的功耗已经在1-10uW之间，每个通道的面积已经在0.005mm2-0.02mm2之间。根据自然杂志的研究报道[6]，脑机接口呈现出类似于“摩尔定律”的发展规律，每隔7至7.5年，脑机接口所能够直接采样的通道数量翻倍，如图8-3所示。以近年的研究成果为例，欧洲的Neuralpixel已实现了超过300通道神经信号采集的脑机接口芯片[7]，而美国的Neuralink公司更是将采集通道数量提升至1024[8]。

图8-3 侵入式脑机接口向着高通量方向持续发展

- 神经信号读出电路设计

- 主要噪声来源

图8-4 神经信号与噪声的关系

侵入式脑机接口系统工作时除了探测LFP、AP等信号外，还受到大量无关信号的干扰。如图8-4所示，主要的干扰信号包括电极直流失调（electrode dc offset，EDO），闪烁噪声（flicker noise，1/f噪声），热噪声以及电力线干扰[9]。直流失调来源于电极通道上电极-组织界面处的电化学反应，不同的通道之间的直流失调电压可达1 mV 到50 mV不等。由于直流失调电压远高于信号强度，很容易导致放大器饱和。闪烁噪声主要分布在低频部分，具有功率谱密度与频率成反比的特征，所以也被称为1/f噪声。热噪声则与1/f噪声不同，它在所有频谱中以相同的形态分布，对信号的影响相对有限。电力线干扰是包括EEG、ECG等多种电生理信号中都存在的一种常见噪声源，是50Hz或60Hz市电对放大器输入端之间形成的共模干扰。植入式脑机接口芯片的一个重要工作是在较低面积和功耗开销下，

表8-1 神经信号读出电路一般性指标

| 输入阻抗 | >100MΩ |
| --- | --- |
| 噪声 | <5-10μVrms |
| 带宽 | 1Hz-10kHz |
| 增益 | > 40 dB |
| 共模抑制比 | >60dB |
| 电源抑制比 | >60dB |
| 直流失调抑制 | 50mV |
| 平均每通道功耗 | <10μW |
| 平均每通道面积 | <0.04mm2 |

放大感兴趣的信号，降低噪声影响，提高信号品质。

在保证所采集的神经信号质量的前提下，为了解决上述问题，需要模拟前端电路满足一些要求。表8-1概括了神经信号读出电路的一般性指标。例如，为了避免芯片发热导致的组织升温，平均每个通道的功耗需尽量保持小于10μW。放大器的带宽需覆盖LFP、AP等所需信号的带宽以满足完整的信号探测要求，等效输入噪声在信号带宽内积分后的均方根应小于神经信号幅度，否则神经信号将淹没在噪声中。为了消除市电以及电源波动的干扰，需要保持较高的共模抑制比和电源抑制比，为了避免放大器饱和，需要具备一定的失调抑制能力。根据不同的脑机接口应用，芯片要求通常会高于上述通用标准。

- 直流失调抑制

在侵入式脑机接口芯片中，电极直流失调信号远高于神经信号本身，极易导致放大器饱和。因此，消除直流失调是神经信号读出电路设计的重要问题之一[10]。最为常用且直接的方法是采用高通滤波，通过设置一个Hz以下高通截止频率，可以对神经信号正常放大，而对低频直流信号抑制，从而避免放大器饱和的问题。

图8-5展示了通过电容耦合方式实现高通滤波的经典电路[11]，其中反馈环路中伪电阻Rf和电容C2共同引入了截止频率为1/(2πRfC2)的高通极点，而中频增益则为电容C1和C2之比。尽管该电路在植入式脑机接口芯片中广泛采用，但是其在使用过程中存在芯片面积较大的问题。例如，为了将截止频率降低至1Hz以下且实现较高增益，即使将C2电容的大小选择为约200fF的最小电容，C1电容的大小仍需在10pF以上。这样会占用较大的芯片面积，限制了芯片的集成度和扩展性。此外，该电路所用的伪电阻阻值对工艺敏感，不易

图8-5经典的交流耦合神经信号放大器

精确控制，这会影响放大器截止频率。

另一种消除直流失调的方法是对输出端信号进行低通滤波，并将滤波后的低频信号通过负反馈回传到输入端，在传递函数中引入一个高通极点，实现高通滤波的功能[12]。这种方法能够自动对失调电压信号进行跟踪，并借助驱动放大器的反相输入来抵消它，达到抑制直流失调的效果。如图8-6所示，假设放大器的直流增益为A，反馈回路的直流增益为B，那么输入电压Vin与输出电压Vout的关系为：

Vout=11A+BVin                                                              (8-1)

图8-6 通过在反馈回路中使用低通滤波器实现直流失调抑制

由于放大倍数A远大于1，因此可认为Vout≈Vin/B。

图8-7（a）中电路通过反馈回路中的模拟积分器实现低通滤波，来去除输入信号中的低频部分。这种设计需要放大器在反馈环路中消耗较多功率，并且其单端配置不适于对低噪声、共模抑制比和电源抑制比要求较高的场景[13]。图8-7（b）中电路的低频和直流信号被反馈至差动差分放大器的第二个差分对，从而消除了信号的低频和直流成分[14]。电路中利用了一个R-C滤波器作

图8-7基于模拟的直流伺服环路

为模拟低通滤波器，为了实现较小的高通极点频率，仍需要较大的电阻电容值。

图8-8基于数字的直流伺服环路

近年来，基于数字直流伺服环路的方式被广泛应用于直流失调抑制[15]。图8-8（a）的电路设计利用反馈中的数字低通滤波器来降低功耗，并使用附加的模拟低通滤波器来降低经过数字路径的直流信号的动态范围，从而放宽DAC分辨率的要求。图8-8（b）利用数字电路实现低通滤波，消除了大电容或者放大器的使用，有效的降低了面积和功耗，且可实现较为精确的极点控制。利用数字反馈的典型直流失调抑制范围约在±50mV左右。

- 斩波技术

为了消除放大器中的低频噪声，可以采用多种方法，包括自调零技术、相关双采样技术和斩波技术等[16]。在低噪声和低功耗的要求下，斩波技术是一种在侵入式脑机接口芯片设计中较为普遍的选择[17-19]。

斩波技术的主要原理是利用周期性方波的交流调制信号，将失调电压以及闪烁噪声等低频信号调制至高频，然后通过低通滤波器来消除这些噪声。其具体工作原理如图8-9所示[16]。输入为Vin的脑电信号在方波信号m1(t)控制的开关切换下完成两个信号在时域内的相乘。由于方波信号本身的频率特性，输入信号将被调制到斩波频率的奇次谐波频率处。经过本次调制后，神经信号、直流失调、低频噪声以及热噪声都将被放大器A放大。输出信号VA即为神经信号和噪声的简单叠加和倍增。VA信号再次在m1(t)控制的开关切换下完成两个信号在时域内的相乘。此时，闪烁和低频噪声如同Vin第一次被调制一样，被调制到了斩波频率的奇次谐波频率处。然而输入信号Vin由于经历了两次调制，其基频处作为二次谐波信号强度最高，最后通过滤波器将高于带宽的干扰信号和谐波分量进行滤波即可获得

图8-9斩波放大器原理

经过放大的神经信号。

图8-10所示为最简单的斩波器设计，通过对反向开关ϕ1和ϕ1的控制，实现了信号的极性反转，该操作相当于对原始信号乘以+1和-1，等效于在时域实现两个信号的乘积。

图8-10斩波器电路

- OTA设计

跨导放大器是低噪声神经放大器中的核心电路[20]。图8-11（a）中电流镜OTA的主极点位于第二级，不需要补偿电容来保持稳定，但是其增益有限，噪声和相位裕度之间也需要折衷。图8-11（b）为带米勒补偿电容的两级运算放大器。图8-11（c）为折叠共源共栅OTA，虽然功耗较高，但可以单级即可提供较高增益。图8-11（d）所示的套筒式OTA可实现较高增益，但是输入范围和电压摆动受到限制。表8-2总结了图8-11中不同 OTA 拓扑

表8-2 常用OTA设计总结

| OTA | 增益 | 噪声 |
| --- | --- | --- |
| 电流镜 | gm1(gm9ro9ro8¦¦ gm10ro10ro6) | (gm1+2gm3+ gm7)16kT/3(gm1)2 |
| 米勒 | gm1(ro2¦¦ro4) gm7ro7 | (gm1+gm3)16kT/3gm1 |
| 折叠式共源共栅 | gm1α[(gm10ro10gm12ro12R) ¦¦ (gm8ro8ro6)] | (gm1+gm6+2/R)16kT/3gm1 |
| 套筒式共源共栅 | gm1(gm4ro3ro4¦¦ gm6ro6ro8) | (gm1+gm8)16kT/3gm1 |

图8-11 常用低噪声OTA电路结构

的增益和输入参考噪声[21]。

- 模数转换器

根据图8-4对脑电信号的分析，可以得知，要完整采集AP及10kHz频率以下的神经信号，单个通道的采样率超过20kHz即可。此外，植入式脑电信号的动态范围通常在10μV至10mV之间，相应的动态范围约为1000。因此，对于信号的数字化，使用10位比特数已经足够。如图8-12所示，模数转换器（analog to digital converter, ADC）主要有四种架构，不同的架构适用于不同的场景。根据上述对采样率和比特数的分析，逐次逼近寄存器模数转换器（successive approximation register analog-to-digital converter，SAR ADC）和Delta-Sigma ADC一般被认为更加适用于侵入式脑机接口芯片中[22, 23]。

图8-13 SAR ADC组成和工作原理

图8-12常用于神经信号模数转化的ADC架构

大部分脑机接口芯片采用的模数转换器结构是SAR。图8-13展示了采样保持电路、电容阵列、比较器以及逻辑电路等SAR ADC的主要组成部分。首先当采样时钟电平为高时，采样保持电路跟随模拟信号，当采样时钟电平为低时，模拟信号被存储在电容阵列中。比较时钟上升沿时，比较器对两端电压进行比较并根据结果输出“1”或者“0”。逻辑电路根据输出结果控制电容阵列开关实现二值搜索，进而使得最终的SAR ADC的数字输出结果逼近采样电压。SAR ADC不需要高增益和高带宽的运算放大器来保持高精度和线性度。逻辑主要由数字电路组成，其速度和功耗可以随着深亚微米CMOS技术的发展而下降。在一般使用电容DAC的情况下，不消耗静态功率，因此功耗与采样率成比例下降。

近年

图8-14 SAR ADC组成和工作原理

来Delta-Sigma ADC逐渐在脑机接口芯片中得到重视[24]。如图8-14所示，它采用了Σ-Δ调制技术，对模拟输入信号进行高频率的过采样并通过模拟积分和差分运算来实现对噪声的滤波和抑制[25]。Σ-Δ调制器的输出是一个高速的1位位流（bitstream）信号，代表了模拟输入信号的模数转换结果。输入模拟量越高，输出的“1”的数量就越多。Σ-Δ调制器的输出位流信号随后会经过数字信号处理模块。该模块通常包括数字滤波器和抽取滤波器。数字滤波器对高速的位流信号进行低通滤波，并将其转换为高精度的多比特信号，同时去除高频噪声和滤除残余的量化噪声。数字滤波器的输出速率与采样率相同，仍然包含了过采样中的大量冗余样本，抽取滤波器对这些样本进行抽取，使其降低至信号原本的频率，并没有导致任何信息的丢失。

- 读出架构设计

随着脑机接口技术的发展，与更多通道读出的需求，出现了不同的多通道架构[26, 27]。经典的结构如图8-15（a）所示，每个通道独占一个模拟前端，多个通道的信号通过模拟多路复用器时分多路复用传输到ADC进行数字化。由于多个通道时分复用，ADC的采样率需随着通道数增加而增加，提高了ADC设计复杂度。一些多通道采样芯片中都采用16或者20个通道时分复用一组ADC的设计，将ADC的采样率控制在320K或者200K左右[28-30]。此外，多通道模拟信号容易受到模拟多路复用器中串扰噪声的影响，因此需要额外的设计策略来提高噪声容限。采用这种架构设计的芯片，电路面积随着通道数的增加几乎线性增长。

（a）

(b)

图8-15多通道读出架构

为了实现更加高通量的神经信号记录，近年来图8-15（b）所示的快速时分复用技术逐渐受到更多的关注[31-34]。该电路通过前置的模拟多路选择器实现对模拟前端放大电路的时分复用。虽然该电路结构通过减少放大器和ADC的数量以降低芯片面积，但是每个通道都有其自己独立的、随时间变化的等效输入直流失调。除此之外，噪声混叠也是一个具有挑战的问题。

近年来，直接采用低功耗高动态范围的ADC对神经信号进行数字化也在侵入式脑机接口芯片中得到实践。这种方法取消了低噪放大器，可以更大程度地减少芯片功耗和面积开销，并通过具有较高采样率ADC所固有的噪声整形能力，实现系统低噪声性能[35, 36]。

- 神经信号处理电路设计

图8-16 利用阈值和NEO算子进行尖峰检测

针对不同的应用场景，可以对数字化的神经信号进行压缩，降低数据传输给系统带宽和功耗带来的压力，也可以在芯片上对神经信号进行直接处理，以进行及时的闭环神经调控[37]。本小节将分别简述神经信号压缩、特征提取以及分类常用算法以及相关电路设计。

- 信号压缩

在侵入式脑机接口芯片上进行数据压缩是一种减少数据传输带来的系统功耗的有效解决方案[28]。对于动作电位记录，一种方法是在芯片上检测并仅传输尖峰信号，从而减少数据量。如图8-16，尖峰检测可以通过在硬件中实现预定义阈值的比较器来实现，这样可以以很小的功耗实现高达上千倍的压缩比。其中，最简单的片上尖峰监测方法如式（8-2）：

a×i=1n(xi-x)2n-1                                                          (8-2)

a×median|x|0.6745                                                  (8-3)

式（8-2）中的阈值由多个采样点的标准差和一个倍数a决定，一般情况下可使用a = 4。另外，也可通过式（8-3）的方式估计一个噪声标准差来设定阈值。此外，尖峰检测还可以采用非线性能量算子（nonlinear energy operator，NEO)[38]，其表达式如式（8-4）所示：

ψxn=x2n-x(n+1)×x(n-1)                  (8-4)

其中，x(n)是n时刻的采样点。当第一项的值较大且x(n-1)、x(n+1)较小时，ψxn的值会相对较大，这个点可以被认为是一个尖峰。通过应用NEO算子后，可以进行一次阈值分割，以完成尖峰检测。通常情况下，所采用的阈值一般为ψxn平均值的比例系数。

基于式（8-2）和（8-3）的尖峰检测在电路上实现简单，但是基于尖峰检测的压缩会丢失大部分原始波形。因此，在使用该方法时需要权衡数据压缩和信息丢失之间的平衡。

图8-17 利用压缩感知进行信号采样和重建的过程

压缩感知（compressive sensing, CS）理论是一种广泛用于神经信号压缩的方法。该理论利用信号的稀疏性，实现亚奈奎斯特采样和接近无损的信号重构[39]。大多数自然和人造信号X∈Rn均可被表示为一个稀疏基矩阵Ψ和稀疏系数s的乘积，即X = Ψs。压缩采样的过程则可用以下方程表示：

Y = ΦX                                                                (8-5)

观测矩阵Φ∈Rm×n是压缩感知中的测量矩阵，其中m < n。通过观测矩阵Φ对X进行采样，我们可以得到较低维度的测量值Y，这样就可以实现减少传输数据量的效果。图8-17展示了信号压缩和重建的完整过程。原始信号X表示为Ψ和s的乘积，即X= Ψs。采样过程由Y = ΦX表示。因此，可以得到Y = ΦΨs = Θs。Θ是一个m×n的矩阵。显然，由于所得的线性方程的数量大于未知数的数量，因此有无限多个s的解。从信号稀疏性出发，我们希望找到一个稀疏向量作为解。为了实现这一点，可以使用L1范数来获得稀疏向量s：

s = argmin||S′||1                                                                                   (8-6)

在植入式芯片设计中，通常只考虑压缩采样过程，由外部计算设备完成信号的重建即可。为了保证数据能够被可靠地恢复，测量矩阵Φ需要满足有限等距性质（restricted isometry property, RIP）。满足该条件的矩阵可以是随机高斯矩阵、随机伯努利矩阵、哈达玛矩阵等。为了降低芯片实现的成本，可将矩阵中元素分配为0、±1、±(1/8)、±(2/8)、...、±(7/8)来近似伯努利矩阵[40, 41]。如图8-18所示，展示了在植入式芯片上实现压缩感知的电路结构图。主要目标是实现式（8-5）中的向量矩阵乘法运算，以实现信号的降维。在模拟实现中，多个通道的信号在模拟域进行采样和求和，最后将压缩后数据流转化为数字信号，在这一过程中，感知矩阵中的元素通过由时序和数字控制的采样或者增益控制实现。

(a)

(b)

图 8-18压缩感知的（a）模拟和（b）数字实现

在数字实现中，通过对数值的移位或者可编程增益放大器的控制，实现采样矩阵。

除了压缩感知方法以外，还有主成分分析（principal component analysis，PCA）[42]等其他方法可进行数据压缩。然而，对于植入式芯片的片上数据压缩来说，这些方法的代价通常过高，难以适用。

- 特征提取

特征提取是脑机接口系统中的关键步骤之一，旨在从神经信号中提取有用的信息，并将其转化为可操作的命令或控制信号。通过在芯片上进行这些特征提取过程，可以降低信号传输和处理的延迟，提高系统的实时性和响应速度，这对于闭环神经调控等应用领域来说尤为关键。神经信号中的特征可以包括频率、时域、相位、振幅等方面的信息，根据应用或研究目标的不同，所选取的特征也有所不同。例如，上节所述尖峰信号在脑机接口的运动解码中是常用特征，相位锁定值（Phase Lock Value，PLV）可以评估两个信号之间的同步性，高频振荡（high frequency oscillations, HFO）特征则与帕金森和癫痫病具有相关性[43, 44]。表8-3列出了常见的神经信号特征，包括频率相关特征、相位相关特征以及时域特征。

表8-3 脑机接口应用中常见特征

| 相位相关特征 | 相位锁定值 |
| --- | --- |
|  | 相位谐振耦合 |
| 频率相关特征 | 高频振荡 |
|  | 跨频耦合 |
|  | 能量谱 |
| 其他 | 线长 |

能量谱是神经信号中最为常用的特征之一，可以用于分析不同频段的活动[45]。如LFP信号中的α波、β波、γ波段的能量谱被普遍认为和帕金森、癫痫等疾病具有相关性。能量谱可利用式（8-7）在离散时间域进行计算：

SE=1Nt=1Ny2                                                                                          (8-7)

其中y为经过带通滤波后的神经信号。带通滤波器可采用有限冲激响应（finite impulse response, FIR）滤波器实现。数字FIR可简单看作对时域信号的卷积计算，如式所示：

yN=i=0N-1bix[i]                                                                                       (8-8)

图 8-19 采用全并行和全串行的FIR滤波器实现

其中x[i]和y[i]分别为输入和输出信号，N-1为FIR滤波器的阶数，bi为滤波器系数。式(8-8)可以通过图8-19中展示的全并行或全串行的实现方法来实现。除此之外，还可以采用部分并行或部分串行的方式来在面积和速度之间进行折衷[46]。由于系数bi是固定值，因此乘法和累加计算存在多种优化方式。例如，可以将固定的系数值分解为移位操作和查找表等。

高频振荡（high frequency oscillations, HFO）是指在某个频率以上的谱能量。例如，高于200Hz的HFO和高于80Hz的HFO被分别认为和帕金森、癫痫的发生和发展过程有一定关系，是疾病诊断和治疗的潜在生物标记物。

线长（line-length，LL）特征是一种在癫痫检测中广泛使用的特征[47]，其计算如式 (8-8)。LL衡量了信号在时间上的快速变化程度。当信号存在高振幅、剧烈变化或高频振荡时LL值将会迅速增加。

LL=1Nt=1N|xt-xt-1|                                                                         (8-9)

线长特征计算简单，目前被很多用于进行癫痫控制的神经调控仪器所采用[48]。

相位锁定值（phase locking value, PLV）是常用于评估神经信号之间相位同步性的特征，在神经科学研究和脑疾病诊断中广泛使用[49]。PLV通过计算信号之间的相位差来评估相位同步性，可采用式（8-9）的方式计算。N为检测时间窗内数据采样点，∆∅i为两个信号之间第i个样本的瞬时相位差。PLV的取值范围在0到1之间，其中0表示没有相位同步，1表示完全的相位同步。

PLV=1N(i=0N-1sin⁡(∆∅i))2+(i=0N-1cos⁡(∆∅i))2                              (8-10)

从时频分析中提取信号瞬时相位常用的方法是采用希尔伯特变换（Hilbert transformation）。图8-19所示的电路结构或其变形结构可以实现希尔伯特变换，从而获得输入信号x的解析信号xa，其中包含了原始实值信号的幅度和相位信息。

xa=Rex+jIm(x)                                                 (8-11)

相位计算可通过式(8-12)计算获得：

∅=Im(x)Re(x)                                                          (8-12)

两个信号之间的瞬时相位差∆∅ = ∅1-∅2。

除了PLV之外，相位谐振耦耦合（Phase-Amplitude Coupling）可用于描述不同频率带的振幅与相位之间的关联关系，其计算方式与PLV类似，其中At为调制幅度包络。

PAC=1N(i=0N-1Atsin⁡(∆∅i))2+(i=0N-1cos⁡At(∆∅i))2                          (8-13)

在侵入式脑机接口芯片中，为了实现反正切函数、开方以及正余弦函数的计算，通常会采用坐标旋转数字计算机（CORDIC）进行近似计算，以降低硬件开销[50]。

- 特征分类

利用特征提取，可得到神经信号的特征向量，而对特征向量进行分类是判断脑疾病发作、解码人脑意图等工作的最后环节。对于侵入式脑机接口芯片而言，除了分类的准确率以外，还需要考虑功耗、面积以及实时性等问题[51]。

支持向量机（support vector machine, SVM）是一种常用的用于片上神经信号特征分类的方法[52, 53]。其数学公式如式（8-14）所示：

fx=i=1NaiKvi,x+b                                           (8-14)

(a)									(b)

图 8-21支持向量机概念以及硬件实现

其中，vi和x分别表示支持向量和特征向量，K代表核函数，ai和b是预先训练得到的参数。特征向量的类别由最终计算值f(x)的符号确定。图8-19展示了SVM分类器的概念和简化硬件实现框图。

在图8-21（a）中，可以看到两个不同类别的支持向量构成了最大化两个类别之间间隔的超平面。对于线性不可分的数据点，核函数可以将其投影到更高的维度，使其变得可分。根据公式（8-14），实现支持向量以及参数的存储和核函数计算是片上完成SVM分类器的关键。图8-21（b）为简化的SVM硬件实现架构，采用存储器对参数和支持向量进行存储，利用状态机进行计算控制。同时，采用并行的核函数计算单元来执行aiKvi,x的计算。与FIR滤波器实现类似，在具体应用中可根据实际需求，采用部分并行或者串行来增减核函数计算模块的数量。

在支持向量机的硬件实现中，核函数的计算至关重要。常用的核函数包括线性函数、多项式函数和径向基函数（radial basis function, RBF）[46, 54, 55]。它们的表达式分别如式（8-14）所示：

Kvi,x=vi∙x

kvi,x=(βvi∙x+γ)M                                        （8-15）

kvi,x=exp(-||vi∙x||22σ)

其中，vi和x是输入向量。β是多项式函数的参数，γ是常数项，M是多项式的阶数，σ是RBF核函数的参数。线性函数或者多项式函数可能会牺牲一部分分类精度，但是所需的硬件资源和计算复杂度较低。RBF函数的计算涉及到复杂的平方、除法和幂运算，对于植入式芯片实现来说是相对复杂，为了简化RBF函数的计算，通常采用CORDIC进行迭代进行近似计算[56-58]。

除了支持向量机，斜决策树（oblique decision tree，ODT）也经常被应用于侵入式的脑机接口芯片中[59, 60]。图8-20（a）展示了一个六分类斜决策树的拓扑结构，由“内部节点”（Node）和“叶节点”（leaf）组成。内部节点通过将超平面和样本特征值进行比较，将样本分配到下一级的内部节点或者叶节点，这一过程不断重复，直到达到叶节点为止。叶节点是决策树的最终输出结果，它给出输入样本的最终分类标签。图8-22（b）为斜决策树二维投影为六个类别的示例。

在斜决策树中，内部节点中超平面和特征向量比较过程可由式（8-16）表示：

f=i=1Naisi+b                                           （8-16）

(a)								(b)

图 8-22决策树概念示意图

其中，N表示输入特征向量维度，si表示特征向量的第i个属性，ai为对应的参数。与支持向量机类似，侵入式芯片仅关注决策树的推理过程，因此主要涉及每个节点的特征值与参数的计算，以及根据计算结果选择下一级节点的逻辑控制过程。决策树的电路实现相对支持向量机来说较为简单，主要涉及加法器、比较器的使用。在脑机接口中，为了进一步提高硬件使用效率和分类准确率，可以通过优化分类的流程或者集成决策树的方法通过将多个决策树组合起来进行分类，提高整体的分类性能[61]。

相较于支持向量机和决策树等传统的机器学习方法，使用深度学习算法可以实现对复杂信号的特征提取和分析，已经被用在癫痫预测、语音解码等诸多领域[62-65]。深度学习算法包含了多种运算，如卷积、池化和激活函数等。相比支持向量机和决策树，深度学习算法的参数量和计算量都大大增加。然而，在侵入式芯片存储器大小受限的情况下，需要将神经网络模型压缩至数十至百kB以下[66]。为了降低算法和硬件实现的难度，很多工作首先对单个通道的数据进行一维卷积，然后再对多个通道的数据特征进行空间方向上的卷积。如图8-23所示，为采用上述方法进行神经信号处理的卷积神经网络示例[67]。通过这样的优化，神经网络的硬件实现变得和FIR滤波器类似，从而实现了在较低硬件开销下分

图 8-23 采用卷积和脉冲神经网络进行神经信号的分类示例

类能力的提升。

除上述方法外，基于神经形态计算（neuromorphic computing）的特征提取和处理在脑机接口应用中也受到了广受关注[68-70]。这种方法受到大脑的启发，使用脉冲信号进行信息编码，并利用神经元-突触模型来实现信息的处理和传递。与传统计算方式相比，神经形态计算具有极高的能量效率[71]，同时由于其本身就是对大脑的模拟，所以其天然地适用于大脑的尖峰信息的处理[72]。例如，已有研究工作利用三层脉冲神经网络对运动皮层的脉冲信号进行解码，实现了猴子脑控玩电脑游戏，以及利用神经形态计算进行癫痫检测[73, 74]。

- 无线能量传输与通讯电路设计

无线能量传输和无线通信电路是植入式芯片中的两个关键电路。无线能量传输电路的主要功能是将能量从外部传输到植入式芯片，以实现芯片的供电或电池充电。通过无线电频率的电磁场耦合或电磁感应方式，外部的能量发射器或充电设备可以将能量传输到植入式芯片上的能量接收器。

图 8-24用于植入式芯片的无线能量传输电路框图

常见的无线能量传输方式包括感性耦合、电磁波辐射、超声波和近红外光等[75]。在脑机接口芯片中，感性耦合是最常用的传输方式[76]。它利用线圈之间的电磁感应原理，通过磁场传输能量。其他方式如电磁波辐射、超声波和近红外光在脑机接口应用中也被使用[77]。图8-24给出了一个典型的无线能量传输电路结构。元件在功率放大器驱动下，将电能转换为时变的电磁场或超声波等，这些能量穿透人体组织后向侵入式的芯片传递。植入式芯片接收到的交流电载波信号经过片上电源管理模块整流和稳压调节处理后，再传递给负载。电源管理模块中的整流器用于将输入交流信号转换为未调节的直流信号，稳压电路进一步抑制电源纹波，以实现对负载的稳定供电。

- 整流电路与稳压电路

图8-25（a）所示为最简单的半波镇流器，由二极管和滤波电容组成。当输入交流电压Vin处于正向阶段且Vin高于输出电压Vrec时，二极管D1导通。在另外半个周期，二极管反向偏置。图8-25（b）所示为植入式芯片中常见的全波桥式整流器，该电路由四个二极管组成一个桥式电路和滤波电容构成。在输入交流电压Vin的正向阶段且两端电压高于Vrec时，D1，D2开启。在交流电压Vin的另外半个周期且两端压差高于Vrec时，D3，D4开启。在全波整流器的每个周期中，Crec被充电两次，因此纹波较半波整流更小。同时由于在每个输入周期内，都能够利用两个二极管进行整流，有效地利用了输入信号的能量，效率更高。在整流电路中，除了使用二极管，还可以使用二极管连接的晶体管，但无论是二极管还是晶体管，它们都会受到正向偏置压降的影响。PN结二极管的压降约为0.7V，而MOS二极管连接压降等于阈值电压。因此，对全波整理电路，输出将降低两倍二极管压降或晶体管阈值电压。

图 8-25 半波整流与全波整流电路

(a)								(b)

图 8-26交叉耦合以及采用比较器的整流电路

图8-26（a）中展示了栅极交叉耦合的整流电路[78]。它由两个NMOS晶体管和两个PMOS晶体管组成，其中，NMOS晶体管的栅极连接到对向晶体管的漏极。每个周期内，其中一个NMOS 晶体管处于线性区域，其阈值效应可以忽略不计，从而将正向压降降低为单个阈值电压。采用8-26（b）所示的比较器控制的主动整流器可以进一步抑制晶体管在开-关转换期间的反向漏电流导通，从而进一步提到电路效率。

- 稳压电路

整流电路的输出通常包含了大量纹波。尽管可以通过增大电容来有效降低波纹，但在植入芯片中，由于内外线圈之间的距离、错位等因素，导致内部线圈获得的交流电压本身就存在显著波动，因此需要进一步对输出电压进行稳压。植入式芯片中往往采用低压差稳压器（low dropout regulator, LDO），来减小功率损耗。LDO属于线性稳压器的一种，它利用一个高增益放大器来控制BJT或者MOSFET传输管的电阻，使得该电阻和负载电阻组成的电阻分压器始终保持所需比例，进而产生预期的输出电压，图8-27给出了典型的LDO电路结构，主要包含了误差放大器（EA）以及其控制的传输管Mp1。很显然，LDO的功率损耗和Vrec-Vout成线性关系，由于Vrec和Vout之差只在数百毫伏之间，因此LDO仍可以保证较高的转换效率。事实上，由于LDO结构相对简单，面积小且噪声较低，在大多数低功耗应用中，线性调节器都采用低压差（LDO）

图 8-27 低压差稳压器基本结构

调节器。

- 低频无线数据传输

图 8-28 典型的窄带发射机电路结构

采用无线数据传输可以实现侵入式脑机接口芯片与其他设备之间的实时数据交换，同时避免有线传输造成的手术创伤和感染等问题，是当前脑机接口系统的发展趋势[79]。因此将数字信号传输到身体外部接收设备，是脑机接口芯片设计中一个非常重要的考虑因素。目前脑机接口芯片中所采用的无线通信技术主要分为窄带通信和超宽带通信两类[80]。图8-28为典型的窄带发射机电路结构框图，主要包含了基带信号处理电路、数模转换器、上变频器以及通信天线四个部分。基带信号信号在压缩、编码等处理后由模数转换器转换为低频模拟信号。上变频器将低频模拟信号与本地振荡器提供的高频载波进行混合后，生成射频调制信号。最后，射频信号通过天线辐射到空间中进行传输。

图 8-29 常用无线通信技术的频带分布

表8-4 常用于侵入式脑机接口的窄带通讯方式比较

| 窄带通讯技术 | 频段 | 通讯范围 | 数据率 | 功耗 |
| --- | --- | --- | --- | --- |
| WMTS | 608-614 MHz, 1395-1400 MHz, 1427-1432 MHz | < 10 m | <2 Mbps | 中：几十mW-1W |
| MICS | 402-405 MHz, 413-419 MHz, 426-432 MHz | < 1m | <2 Mbps | 较低：几mW到数百mW |
| NFC | 13.56 MHz, | < 10 cm | < 10 Mbps | 极低：几毫瓦到几十mW |
| Bluetooth | 2.4 GHz | < 10 m | < 10 Mbps | 低：几mW到几十mW |
| WiFi | 0.9 GHz, 2.4 GHz, 5 GHz | < 100 m | 10Mbps – 1Gbps | 高：几百mW到数W |
| ZigBee | 2.4 GHz | < 100 m | <250 kbps | 较低：几十mW到数百mW |

窄带通讯技术包括医疗领域专门制定的WMTS（Wireless Medical Telemetry Service）、MICS（Medical Implant Communication Service）和蓝牙、WiFi、ZigBee、近场通信等通用通讯标准。WMTS的频段通常在608-614 MHz、1395-1400 MHz和1427-1432 MHz，可用于患者心电图、血氧、血糖等生命体征的数据传输，具有稳定、可靠的特点，但是带宽一般较低。MICS是专门为医疗植入设备而设计，频段一般在402-405 MHz、413-419 MHz和426-432 MHz的通信方式，在植入式传感器、心脏起搏器、神经刺激器等植入式医疗设备中均有应用，其特点是低功耗和高可靠性，其带宽同样较低[81]。ISM（Industrial Scientific Medical）频段中的13.56MHz也常被用于侵入式芯片的近场数据通讯（near field communication），其特点是功率传输可一同在相同的载波频率上实现，但是数据传输率仍然有限[82, 83]。表8-4对上述通讯方式的数据率、传输距离和功耗等方式进行了总结，图8-29为频带分布图。

(a)

(b)

图 8-30 无载波与有载波UWB电路结构

随着侵入式脑机接口系统不断向着高通量方向发展，脑机接口芯片需要在极低的功耗下提供更高的数据带宽。与上述窄带技术不同，带宽是载波频率的0.2倍或者绝对带宽大于500MHz 的信号定义为超宽带（ultra-wideband, UWB）信号[84]，其中0~960  MHz和3.1~10.6 GHz两个频段可用于超宽带通信且要求信号辐射功率谱密度小于-41.3dBm/MHz（图8-29）。超宽带信道容量大，能够实现高通量数据的无线传输，且其固有的低功率辐射要求缓解了功率放大器功耗较高的问题，在最新的脑机接口芯片中得到了广泛应用[85-87]。从超宽带的频谱要求出发，一般在时域中使用的主要是高斯包络、三角包络、半余弦和矩形四种波形，可采用不同的电路设计方法产生上述波形。图8-30分别展示了侵入式脑机接口芯片中采用的典型UWB电路结构，可分为无载波和有载波型[88-90]。图8-30（a）中的无载波UWB发射器在利用多个延迟单元和组合逻辑产生高阶高斯脉冲波形，最后通过功率放大器驱动天线。该电路结构简单，主要以数字电路为主，不需要混频器和本地振荡器，功率损耗较低，但是多路精准的亚纳秒级数字延迟单元在电路与版图上实现较为困难，容易受到噪声和失配的影响。图8-30（b）中所示的基于载波的结构，包括一个脉冲发生器、压控振荡器、一个脉冲整形器以及功率放大器。该电路首先根据数字信号产生窄脉冲控制信号，该控制信号随后触发压控振荡器产生多周期的脉冲包络信号，最后经过脉冲整形之后输出满足要求的信号。由于其工作于触发状态，功耗较低，但是需要振荡器能够快速开启和关闭以避免额外频率的引入。

- 刺激电路

使用电刺激的方式实现对神经元兴奋或抑制的调控，是脑疾病治疗中最常见且重要的功能之一[91]。这种功能被广泛应用于脑机接口设备中，用于治疗各种脑部疾病和恢复功能。例如，深部脑刺激（deep brain stimulation，DBS）就被应用在帕金森的治疗上[92]，而视皮层刺激被证明可以帮助盲人实现部分视觉恢复[93, 94]。

- 常用刺激波形

根据刺激方式的不同，植入式电刺激器主要分为电压控制输出和电流控制输出两种方式。在电压控制的刺激器中，用于产生电刺激的两个电极之间存在刺激电压，刺激电路的输出受到组织和电极之间的阻抗影响。然而，电极与组织之间的阻抗模型是不稳定的，并且会随着刺激过程而变化。因此，在单个刺激周期内，注入到组织中的电荷无法均匀释放。当剩余电荷积累到阈值时，可能会对神经组织造成无法预测的损害。在电流控制刺激中，注入组织的电荷可以由电流模式数模转换器 (DAC) 控制。 通过设置刺激器的振幅和时间，可以精确的控制电荷的注入情况。电流控制输出不受组织负载变化的影响，可对输出到组织中的电荷量进行监控，较电压控制输出而言安全性更加在实际应用中。因此，电流控制输出更为常见。

图8-31 不同刺激方式示意图

如图8-31所示，根据刺激方式的不同，刺激方式可以分为单相（monophasic）和双相刺激（biphasic）。在单相刺激中，电荷的积累可能会导致组织损伤并对影响电极使用寿命。因此，更常用的是双相电流刺激，通过控制电流刺激方向，可以补偿大部分电荷，避免组织损伤和电极损坏的发生[95]。典型的双相刺激波形具备可调的电流强度、脉冲宽度以及刺激间隔。根据刺激脑位置和电极的不同，电流强度一般在10uA至1mA，刺激电压则在2V至10V之间。

图8-32 单极和双极电神经刺激器的简化示意图

在双相刺激中，可采用如图8-32所示的单极（monopolar）或双极（bipolar）电极配置方式。单极刺激使用单个工作电极，利用一对电流源和电流沉产生正相和负相两个阶段的刺激。在双极刺激中，正相和负相两个阶段的刺激由桥式电路控制电流源或电流沉电流在一对电极间的流动方向来产生。在高密度电极阵列中，常采用单极刺激方式，而双极刺激则具有更好的定向刺激效果。

- 神经刺激电路

图8-33典型神经刺激器框图

在脑机接口中，刺激器电路主要包括控制器、数模转换器（DAC）、电流驱动器和升压电路等部分。神经刺激器的输出电流通常由微控制器单元（MCU）或状态机等数字控制模块控制。数字模拟转换器（DAC）将数字信号转换为参考模拟信号，供后续的电流驱动器使用。DAC的输出位数决定了神经刺激器输出电流的精度，一般根据刺激电流的动态范围和跨度进行设计。电流驱动器用于将DAC输出的参考模拟信号转换为刺激器所需的电流输出。如图8-33中所示，神经刺激电路中可采用一对源-沉电流驱动对或者带有H桥的单一电流驱动器来完成单极或双极刺激。电流驱动器需要具有较高的输出阻抗以及高电压兼容性。高输出阻抗可确保刺激器能够有效地将电流传输到不同的负载阻抗。高电压兼容性（high voltage compliance）可保证刺激器在传递所需电流的同时，能够适应负载阻抗的变化，不会导致输出电压达到其限制，确保稳定的电流输出。由于深亚微米工艺的影响，植入式芯片的电压往往无法满足神经刺激的要求，因此需要升压电路来提供系统所需的高电压。

图8-34 6位数模转换器电路示意图

神经刺激电流中常采用的是基于电流源的数模转换器。图8-34展示了一个6位的DAC电路，图中的B0-B5为数字控制码，利用晶体管的宽长比来获得不同大小的电流。通过上拉信号U、下拉信号D以及数字控制码的控制电流方向和大小，电极上可产生所需的刺激电流，Vbp和Vbn为数模转换器的偏置电压，用于产生基准电流[96]。

图8-35 电流模式数模转换器控制电流镜用于神经刺激

电流微刺激

通过电流镜配合DAC输出是刺激器电路中常用的方法。图8-35中展示了利用单管MOS电流镜、共源共栅电流镜以及宽摆幅电流镜的神经刺激电路。单管电流镜的主要问题是输出阻抗较低，采用共源共栅或宽摆幅电流镜虽然较大程度的提高了输出阻抗，但是摆幅均存在不同程度的降低，影响高电压兼容性[97]。除此之外，还可以采用电压-电流转换

图8-36 通过电压-电流转换实现神经电刺激

的方式产生所需刺激电流。图8-36中，利用跨导放大器负反馈可以形成一个简单的电流驱动器，电压型DAC控制Vref或者控制处于线性区的可变电阻实现对刺激电流的控制[98]。采用图8-36的方式可显著提高输出阻抗，但是输出电流线性度易受到线性区工作晶体管的影响。

电荷泵电路被广泛应用于刺激器电路中，以产生高直流供电电压并向组织输送所需的足够电流[99]。图8-37所示是一个N级电荷泵电路的示意图，主要包括电容、电荷传输开关（charge transfer switch，CTS）。其中每个第i级由一个CTS和电容C组成，最后一对CTS和电容CL构成了输出级[100]。电荷泵利用周期性信号CLK与CLK_B的控制来实现对电容的充电，间接地将电荷从一个电容器传递至另一个电容并在过程中不断积累电荷，从而使得输出电压升高。CTS的主要作用是将电荷不可逆的从输入传输至输出，CTS的设计是不同类型电荷泵的主要差异所在[101]。

图8-37 通过电压-电流转换实现神经电刺激

- 电荷平衡

在现实应用中，确保双相刺激中正负两相的总电荷量相等，以实现刺激电荷的平衡非常重要。尽管可以通过时钟来控制两相刺激，但如果刺激过程没有进行校准，阳极和阴极电流之间仍然存在一定误差。如果刺激器在短时间内进行大量的高负荷刺激，这种误差将会被迅速放大，导致大量的电荷积累。为了避免这种情况，通常在刺激电极上串联一个隔离电容器，以限制总的刺激电荷量[102]。为了保证刺激器的输出电压范围，隔离电容值一般较大。在功能性电刺激中，隔离电容器的容量通常在几十纳法到几微法之间，物理尺寸较大，难以在芯片上集成[98]。

除了采用隔离电容方法外，还可以在刺激过程中对两相刺激电流进行匹配和校准。如果电流匹配不足以实现净零电荷要求，可使用额外的放电阶段来去除残余电荷。被动放电直接简单地将刺激电极接地或连接到公共或参考电极。被动放电的缺点是放电电流取决于负载和电极阻抗，无法很好地控制。如果阻抗太低，放电电流容易过高，有损坏组织的可能，因此需要额外的电流限制电路；如果阻抗太高，放电时间久，残余电荷可能无法在下次刺激之前完全清除。

# 参考文献

[1]	U. Chaudhary, N. Birbaumer, and A. Ramos-Murguialday, "Brain–computer interfaces for communication and rehabilitation," Nature Reviews Neurology, vol. 12, no. 9, pp. 513-525, 2016.

[2]	L. Kuhlmann, K. Lehnertz, M. P. Richardson, B. Schelter, and H. P. Zaveri, "Seizure prediction—ready for a new era," Nature Reviews Neurology, vol. 14, no. 10, pp. 618-630, 2018.

[3]	G. Schalk and E. C. Leuthardt, "Brain-computer interfaces using electrocorticographic signals," IEEE reviews in biomedical engineering, vol. 4, pp. 140-154, 2011.

[4]	E. C. Leuthardt, G. Schalk, J. Roland, A. Rouse, and D. W. Moran, "Evolution of brain-computer interfaces: going beyond classic motor physiology," Neurosurgical focus, vol. 27, no. 1, p. E4, 2009.

[5]	H. Wu, J. Chen, X. Liu, W. Zou, J. Yang, and M. Sawan, "An Energy-Efficient Small-Area Configurable Analog Front-End Interface for Diverse Biosignals Recording," IEEE Transactions on Biomedical Circuits and Systems, 2023.

[6]	I. H. Stevenson and K. P. Kording, "How advances in neural recording affect data analysis," Nature neuroscience, vol. 14, no. 2, pp. 139-142, 2011.

[7]	C. M. Lopez et al., "A Neural Probe with Up to 966 Electrodes and Up to 384 Configurable Channels in 0.13um SOI CMOS," IEEE transactions on biomedical circuits and systems, vol. 11, no. 3, pp. 510-522, 2017.

[8]	E. Musk, "An integrated brain-machine interface platform with thousands of channels," Journal of medical Internet research, vol. 21, no. 10, p. e16194, 2019.

[9]	C. M. Lopez and X. Huang, "Circuits and Architectures for Neural Recording Interfaces," in Biomedical Electronics, Noise Shaping ADCs, and Frequency References: Advances in Analog Circuit Design 2022: Springer, 2023, pp. 45-57.

[10]	A. Bagheri, M. T. Salam, J. L. P. Velazquez, and R. Genov, "Low-frequency noise and offset rejection in DC-coupled neural amplifiers: A review and digitally-assisted design tutorial," IEEE transactions on biomedical circuits and systems, vol. 11, no. 1, pp. 161-176, 2016.

[11]	R. R. Harrison and C. Charles, "A low-power low-noise CMOS amplifier for neural recording applications," IEEE Journal of solid-state circuits, vol. 38, no. 6, pp. 958-965, 2003.

[12]	J. Xu, S. Mitra, C. Van Hoof, R. F. Yazicioglu, and K. A. Makinwa, "Active electrodes for wearable EEG acquisition: Review and electronics design methodology," IEEE reviews in biomedical engineering, vol. 10, pp. 187-198, 2017.

[13]	B. Gosselin, M. Sawan, and C. A. Chapman, "A low-power integrated bioamplifier with active low-frequency suppression," IEEE Transactions on biomedical circuits and systems, vol. 1, no. 3, pp. 184-192, 2007.

[14]	K. A. Ng and P. K. Chan, "A CMOS analog front-end IC for portable EEG/ECG monitoring applications," IEEE Transactions on Circuits and Systems I: Regular Papers, vol. 52, no. 11, pp. 2335-2347, 2005.

[15]	R. Muller, S. Gambini, and J. M. Rabaey, "A 0.013${\hbox {mm}}^{2} $, 5$\mu\hbox {W} $, DC-Coupled Neural Signal Acquisition IC With 0.5 V Supply," IEEE Journal of Solid-State Circuits, vol. 47, no. 1, pp. 232-243, 2011.

[16]	C. C. Enz and G. C. Temes, "Circuit techniques for reducing the effects of op-amp imperfections: autozeroing, correlated double sampling, and chopper stabilization," Proceedings of the IEEE, vol. 84, no. 11, pp. 1584-1614, 1996.

[17]	T. Denison, K. Consoer, W. Santa, A.-T. Avestruz, J. Cooley, and A. Kelly, "A 2 $\mu\hbox{W}$ 100 nV/rtHz Chopper-Stabilized Instrumentation Amplifier for Chronic Measurement of Neural Field Potentials," IEEE Journal of Solid-State Circuits, vol. 42, no. 12, pp. 2934-2945, 2007, doi: 10.1109/jssc.2007.908664.

[18]	C.-Y. Wu, C.-H. Cheng, and Z.-X. Chen, "A 16-channel CMOS chopper-stabilized analog front-end ECoG acquisition circuit for a closed-loop epileptic seizure control system," IEEE transactions on biomedical circuits and systems, vol. 12, no. 3, pp. 543-553, 2018.

[19]	E. Greenwald et al., "A Bidirectional Neural Interface IC With Chopper Stabilized BioADC Array and Charge Balanced Stimulator," IEEE Trans Biomed Circuits Syst, vol. 10, no. 5, pp. 990-1002, Oct 2016, doi: 10.1109/TBCAS.2016.2614845.

[20]	B. Gosselin, "Recent advances in neural recording microsystems," Sensors, vol. 11, pp. 4572-4597, 2011.

[21]	B. Razavi, Design of analog CMOS integrated circuits. 清华大学出版社有限公司, 2005.

[22]	M. S. Chae, W. Liu, and M. Sivaprakasam, "Design optimization for integrated neural recording systems," IEEE Journal of Solid-State Circuits, vol. 43, no. 9, pp. 1931-1939, 2008.

[23]	D. De Dorigo et al., "Fully immersible subcortical neural probes with modular architecture and a delta-sigma ADC integrated under each electrode for parallel readout of 144 recording sites," IEEE Journal of Solid-State Circuits, vol. 53, no. 11, pp. 3111-3125, 2018.

[24]	H. Chandrakumar and D. Marković, "A 15.2-ENOB 5-kHz BW 4.5-$\mu $ W Chopped CT $\Delta\Sigma $-ADC for Artifact-Tolerant Neural Recording Front Ends," IEEE Journal of Solid-State Circuits, vol. 53, no. 12, pp. 3470-3483, 2018.

[25]	S. Pavan, R. Schreier, and G. C. Temes, Understanding delta-sigma data converters. John Wiley & Sons, 2017.

[26]	J. Chen, M. Tarkhan, H. Wu, F. H. Noshahr, J. Yang, and M. Sawan, "Recent trends and future prospects of neural recording circuits and systems: A tutorial brief," IEEE Transactions on Circuits and Systems II: Express Briefs, vol. 69, no. 6, pp. 2654-2660, 2022.

[27]	F. Hashemi Noshahr, M. Nabavi, and M. Sawan, "Multi-channel neural recording implants: A review," sensors, vol. 20, no. 3, p. 904, 2020.

[28]	 D.-Y. Yoon, S. Pinto, S. Chung, P. Merolla, T.-W. Koh, and D. Seo, "A 1024-channel simultaneous recording neural SoC with stimulation and real-time spike detection," in 2021 Symposium on VLSI Circuits, 2021: IEEE, pp. 1-2.

[29]	C. M. Lopez et al., "An implantable 455-active-electrode 52-channel CMOS neural probe," IEEE Journal of Solid-State Circuits, vol. 49, no. 1, pp. 248-261, 2013.

[30]	S.-Y. Park et al., "A miniaturized 256-channel neural recording interface with area-efficient hybrid integration of flexible probes and CMOS integrated circuits," IEEE Transactions on Biomedical Engineering, vol. 69, no. 1, pp. 334-346, 2021.

[31]	N. S. K. Fathy, J. Huang, and P. P. Mercier, "A digitally assisted multiplexed neural recording system with dynamic electrode offset cancellation via an LMS interference-canceling filter," IEEE Journal of Solid-State Circuits, vol. 57, no. 3, pp. 953-964, 2021.

[32]	J. P. Uehlin, W. A. Smith, V. R. Pamula, S. I. Perlmutter, J. C. Rudell, and V. S. Sathe, "A 0.0023 mm $^ 2$/ch. Delta-Encoded, Time-Division Multiplexed Mixed-Signal ECoG Recording Architecture With Stimulus Artifact Suppression," IEEE transactions on biomedical circuits and systems, vol. 14, no. 2, pp. 319-331, 2019.

[33]	N. Pérez-Prieto, Á. Rodríguez-Vázquez, M. Álvarez-Dolado, and M. Delgado-Restituto, "A 32-channel time-multiplexed artifact-aware neural recording system," IEEE Transactions on Biomedical Circuits and Systems, vol. 15, no. 5, pp. 960-977, 2021.

[34]	M. Sharma, H. J. Strathman, and R. M. Walker, "Verification of a rapidly multiplexed circuit for scalable action potential recording," IEEE transactions on biomedical circuits and systems, vol. 13, no. 6, pp. 1655-1663, 2019.

[35]	C. Kim, S. Joshi, H. Courellis, J. Wang, C. Miller, and G. Cauwenberghs, "Sub-$\mu $ V rms-Noise Sub-$\mu $ W/Channel ADC-Direct Neural Recording With 200-mV/ms Transient Recovery Through Predictive Digital Autoranging," IEEE Journal of Solid-State Circuits, vol. 53, no. 11, pp. 3101-3110, 2018.

[36]	H. Kassiri et al., "Rail-to-rail-input dual-radio 64-channel closed-loop neurostimulator," IEEE Journal of Solid-State Circuits, vol. 52, no. 11, pp. 2793-2810, 2017.

[37]	S. Zhao, C. Fang, J. Yang, and M. Sawan, "Emerging energy-efficient biosignal-dedicated circuit techniques: A tutorial brief," IEEE Transactions on Circuits and Systems II: Express Briefs, vol. 69, no. 6, pp. 2592-2597, 2022.

[38]	S. Gibson, J. W. Judy, and D. Marković, "Spike sorting: The first step in decoding the brain: The first step in decoding the brain," IEEE Signal processing magazine, vol. 29, no. 1, pp. 124-143, 2011.

[39]	R. G. Baraniuk, "Compressive sensing [lecture notes]," IEEE signal processing magazine, vol. 24, no. 4, pp. 118-121, 2007.

[40]	M. Shoaran, M. H. Kamal, C. Pollo, P. Vandergheynst, and A. Schmid, "Compact low-power cortical recording architecture for compressive multichannel data acquisition," IEEE transactions on biomedical circuits and systems, vol. 8, no. 6, pp. 857-870, 2014.

[41]	X. Liu et al., "A fully integrated wireless compressed sensing neural signal acquisition system for chronic recording and brain machine interface," IEEE Transactions on biomedical circuits and systems, vol. 10, no. 4, pp. 874-883, 2016.

[42]	T. Wu, W. Zhao, H. Guo, H. H. Lim, and Z. Yang, "A streaming PCA VLSI chip for neural data compression," IEEE transactions on biomedical circuits and systems, vol. 11, no. 6, pp. 1290-1302, 2017.

[43]	T. E. Özkurt et al., "High frequency oscillations in the subthalamic nucleus: a neurophysiological marker of the motor state in Parkinson's disease," Experimental neurology, vol. 229, no. 2, pp. 324-331, 2011.

[44]	J. Jacobs et al., "High-frequency oscillations (HFOs) in clinical epilepsy," Progress in neurobiology, vol. 98, no. 3, pp. 302-315, 2012.

[45]	M. X. Cohen, Analyzing neural time series data: theory and practice. MIT press, 2014.

[46]	J. Yoo, L. Yan, D. El-Damak, M. A. B. Altaf, A. H. Shoeb, and A. P. Chandrakasan, "An 8-channel scalable EEG acquisition SoC with patient-specific seizure classification and recording processor," IEEE journal of solid-state circuits, vol. 48, no. 1, pp. 214-228, 2012.

[47]	L. Guo, D. Rivero, J. Dorado, J. R. Rabunal, and A. Pazos, "Automatic epileptic seizure detection in EEGs based on line length feature and artificial neural networks," Journal of neuroscience methods, vol. 191, no. 1, pp. 101-109, 2010.

[48]	R. B. Duckrow and T. K. Tcheng, "Daily variation in an intracranial EEG feature in humans detected by a responsive neurostimulator system," Epilepsia, vol. 48, no. 8, pp. 1614-1620, 2007.

[49]	F. Mormann, K. Lehnertz, P. David, and C. E. Elger, "Mean phase coherence as a measure for phase synchronization and its application to the EEG of epilepsy patients," Physica D: Nonlinear Phenomena, vol. 144, no. 3-4, pp. 358-369, 2000.

[50]	K. Abdelhalim, H. M. Jafari, L. Kokarovtseva, J. L. P. Velazquez, and R. Genov, "64-Channel UWB Wireless Neural Vector Analyzer SOC With a Closed-Loop Phase Synchrony-Triggered Neurostimulator," IEEE Journal of Solid-State Circuits, vol. 48, no. 10, pp. 2494-2510, 2013, doi: 10.1109/jssc.2013.2272952.

[51]	Y. Wei et al., "A review of algorithm & hardware design for AI-based biomedical applications," IEEE transactions on biomedical circuits and systems, vol. 14, no. 2, pp. 145-163, 2020.

[52]	S.-A. Huang, K.-C. Chang, H.-H. Liou, and C.-H. Yang, "A 1.9-mW SVM processor with on-chip active learning for epileptic seizure control," IEEE Journal of Solid-State Circuits, vol. 55, no. 2, pp. 452-464, 2019.

[53]	L. Feng, Z. Li, and Y. Wang, "VLSI design of SVM-based seizure detection system with on-chip learning capability," IEEE transactions on biomedical circuits and systems, vol. 12, no. 1, pp. 171-181, 2017.

[54]	M. A. B. Altaf, C. Zhang, and J. Yoo, "A 16-channel patient-specific seizure onset and termination detection SoC with impedance-adaptive transcranial electrical stimulator," IEEE Journal of Solid-State Circuits, vol. 50, no. 11, pp. 2728-2740, 2015.

[55]	M. A. B. Altaf and J. Yoo, "A 1.83$\mu $ J/classification, 8-channel, patient-specific epileptic seizure classification SoC using a non-linear support vector machine," IEEE Transactions on Biomedical Circuits and Systems, vol. 10, no. 1, pp. 49-60, 2015.

[56]	K. Abdelhalim, V. Smolyakov, and R. Genov, "Phase-synchronization early epileptic seizure detector VLSI architecture," IEEE transactions on biomedical circuits and systems, vol. 5, no. 5, pp. 430-438, 2011.

[57]	M. Shoaib, K. H. Lee, N. K. Jha, and N. Verma, "A 0.6–107 µW energy-scalable processor for directly analyzing compressively-sensed EEG," IEEE Transactions on Circuits and Systems I: Regular Papers, vol. 61, no. 4, pp. 1105-1118, 2014.

[58]	K. H. Lee and N. Verma, "A low-power processor with configurable embedded machine-learning accelerators for high-order and adaptive analysis of medical-sensor signals," IEEE Journal of Solid-State Circuits, vol. 48, no. 7, pp. 1625-1637, 2013.

[59]	M. Shoaran, B. A. Haghi, M. Taghavi, M. Farivar, and A. Emami-Neyestanak, "Energy-efficient classification for resource-constrained biomedical applications," IEEE Journal on Emerging and Selected Topics in Circuits and Systems, vol. 8, no. 4, pp. 693-707, 2018.

[60]	U. Shin et al., "NeuralTree: A 256-Channel 0.227-μJ/Class Versatile Neural Activity Classification and Closed-Loop Neuromodulation SoC," IEEE Journal of Solid-State Circuits, vol. 57, no. 11, pp. 3243-3257, 2022.

[61]	B. Zhu, M. Farivar, and M. Shoaran, "Resot: Resource-efficient oblique trees for neural signal classification," IEEE Transactions on Biomedical Circuits and Systems, vol. 14, no. 4, pp. 692-704, 2020.

[62]	J. Yang and M. Sawan, "From seizure detection to smart and fully embedded seizure prediction engine: A review," IEEE Transactions on Biomedical Circuits and Systems, vol. 14, no. 5, pp. 1008-1023, 2020.

[63]	G. K. Anumanchipalli, J. Chartier, and E. F. Chang, "Speech synthesis from neural decoding of spoken sentences," Nature, vol. 568, no. 7753, pp. 493-498, 2019.

[64]	 J. Liu et al., "4.5 BioAIP: A Reconfigurable Biomedical AI Processor with Adaptive Learning for Versatile Intelligent Health Monitoring," in 2021 IEEE International Solid-State Circuits Conference (ISSCC), 2021, vol. 64: IEEE, pp. 62-64.

[65]	S. Zhao et al., "A 0.99-to-4.38 uJ/class Event-Driven Hybrid Neural Network Processor for Full-Spectrum Neural Signal Analyses," IEEE Transactions on Biomedical Circuits and Systems, 2023.

[66]	V. J. Lawhern, A. J. Solon, N. R. Waytowich, S. M. Gordon, C. P. Hung, and B. J. Lance, "EEGNet: a compact convolutional neural network for EEG-based brain–computer interfaces," Journal of neural engineering, vol. 15, no. 5, p. 056013, 2018.

[67]	S. Zhao, J. Yang, and M. Sawan, "Energy-efficient neural network for epileptic seizure prediction," IEEE Transactions on Biomedical Engineering, vol. 69, no. 1, pp. 401-411, 2021.

[68]	F. Tian, J. Yang, S. Zhao, and M. Sawan, "NeuroCARE: A generic neuromorphic edge computing framework for healthcare applications," Frontiers in Neuroscience, vol. 17, p. 1093865, 2023.

[69]	C. Wang, J. Yang, and M. Sawan, "NeuroSEE: A Neuromorphic Energy-Efficient Processing Framework for Visual Prostheses," IEEE Journal of Biomedical and Health Informatics, vol. 26, no. 8, pp. 4132-4141, 2022.

[70]	C. Wang, C. Fang, Y. Zou, J. Yang, and M. Sawan, "SpikeSEE: An energy-efficient dynamic scenes processing framework for retinal prostheses," Neural Networks, vol. 164, pp. 357-368, 2023.

[71]	C. Fang, C. Wang, S. Zhao, F. Tian, J. Yang, and M. Sawan, "A 510μW 0.738-mm2 6.2-pJ/SOP Online Learning Multi-Topology SNN Processor with Unified Computation Engine in 40-nm CMOS," IEEE Transactions on Biomedical Circuits and Systems, 2023.

[72]	J. Yoo and M. Shoaran, "Neural interface systems with on-device computing: Machine learning and neuromorphic architectures," Current opinion in biotechnology, vol. 72, pp. 95-101, 2021.

[73]	Y. Yang, J. K. Eshraghian, N. D. Truong, A. Nikpour, and O. Kavehei, "Neuromorphic deep spiking neural networks for seizure detection," Neuromorphic Computing and Engineering, vol. 3, no. 1, p. 014010, 2023.

[74]	 F. Tian, J. Yang, S. Zhao, and M. Sawan, "A new neuromorphic computing approach for epileptic seizure prediction," in 2021 IEEE International Symposium on Circuits and Systems (ISCAS), 2021: IEEE, pp. 1-5.

[75]	G. L. Barbruni, P. M. Ros, D. Demarchi, S. Carrara, and D. Ghezzi, "Miniaturised wireless power transfer systems for neurostimulation: A review," IEEE Transactions on Biomedical Circuits and Systems, vol. 14, no. 6, pp. 1160-1178, 2020.

[76]	M. Sawan, Y. Hu, and J. Coulombe, "Wireless smart implants dedicated to multichannel monitoring and microstimulation," IEEE Circuits and systems magazine, vol. 5, no. 1, pp. 21-39, 2005.

[77]	R. Muller, M. M. Ghanbari, and A. Zhou, "Miniaturized wireless neural interfaces: A tutorial," IEEE Solid-State Circuits Magazine, vol. 13, no. 4, pp. 88-97, 2021.

[78]	Y. Lu and W.-H. Ki, "CMOS integrated circuit design for wireless power transfer," 2018.

[79]	S. M. Won, L. Cai, P. Gutruf, and J. A. Rogers, "Wireless and battery-free technologies for neuroengineering," Nature Biomedical Engineering, vol. 7, no. 4, pp. 405-423, 2023.

[80]	M. Sawan, Handbook of Biochips: Integrated Circuits and Systems for Biology and Medicine. Springer, 2015.

[81]	M. N. Islam and M. R. Yuce, "Review of medical implant communication system (MICS) band and network," Ict Express, vol. 2, no. 4, pp. 188-194, 2016.

[82]	G. Simard, M. Sawan, and D. Massicotte, "High-speed OQPSK and efficient power transfer through inductive link for biomedical implants," IEEE Transactions on biomedical circuits and systems, vol. 4, no. 3, pp. 192-200, 2010.

[83]	F. Mounaim and M. Sawan, "Integrated high-voltage inductive power and data-recovery front end dedicated to implantable devices," IEEE Transactions on Biomedical Circuits and Systems, vol. 5, no. 3, pp. 283-291, 2011.

[84]	L. Yang and G. B. Giannakis, "Ultra-wideband communications: an idea whose time has come," IEEE signal processing magazine, vol. 21, no. 6, pp. 26-54, 2004.

[85]	H. Ando, K. Takizawa, T. Yoshida, K. Matsushita, M. Hirata, and T. Suzuki, "Wireless multichannel neural recording with a 128-Mbps UWB transmitter for an implantable brain-machine interfaces," IEEE transactions on biomedical circuits and systems, vol. 10, no. 6, pp. 1068-1078, 2016.

[86]	M. S. Chae, Z. Yang, M. R. Yuce, L. Hoang, and W. Liu, "A 128-channel 6 mW wireless neural recording IC with spike feature extraction and UWB transmitter," IEEE transactions on neural systems and rehabilitation engineering, vol. 17, no. 4, pp. 312-321, 2009.

[87]	 M. Song et al., "A 1.66 Gb/s and 5.8 pJ/b transcutaneous IR-UWB telemetry system with hybrid impulse modulation for intracortical brain-computer interfaces," in 2022 IEEE International Solid-State Circuits Conference (ISSCC), 2022, vol. 65: IEEE, pp. 394-396.

[88]	 G. Lee, S. Lee, J.-H. Kim, and T. W. Kim, "21.1 A 1.125 Gb/s 28mW 2m-Radio-Range IR-UWB CMOS Transceiver," in 2021 IEEE International Solid-State Circuits Conference (ISSCC), 2021, vol. 64: IEEE, pp. 302-304.

[89]	N.-S. Kim and J. M. Rabaey, "A high data-rate energy-efficient triple-channel UWB-based cognitive radio," IEEE Journal of Solid-State Circuits, vol. 51, no. 4, pp. 809-820, 2016.

[90]	 E. Allebes et al., "21.2 A 3-to-10GHz 180pJ/b IEEE802. 15.4 z/4a IR-UWB coherent polar transmitter in 28nm CMOS with asynchronous amplitude pulse-shaping and injection-locked phase modulation," in 2021 IEEE International Solid-State Circuits Conference (ISSCC), 2021, vol. 64: IEEE, pp. 304-306.

[91]	S. M. Won, E. Song, J. T. Reeder, and J. A. Rogers, "Emerging modalities and implantable technologies for neuromodulation," Cell, vol. 181, no. 1, pp. 115-135, 2020.

[92]	P. Limousin and T. Foltynie, "Long-term outcomes of deep brain stimulation in Parkinson disease," Nature Reviews Neurology, vol. 15, no. 4, pp. 234-242, 2019.

[93]	A. N. Foroushani, C. C. Pack, and M. Sawan, "Cortical visual prostheses: from microstimulation to functional percept," Journal of neural engineering, vol. 15, no. 2, p. 021005, 2018.

[94]	J. Coulombe, M. Sawan, and J.-F. Gervais, "A highly flexible system for microstimulation of the visual cortex: Design and implementation," IEEE transactions on biomedical circuits and systems, vol. 1, no. 4, pp. 258-269, 2007.

[95]	D. R. Merrill, M. Bikson, and J. G. Jefferys, "Electrical stimulation of excitable tissue: design of efficacious and safe protocols," Journal of neuroscience methods, vol. 141, no. 2, pp. 171-198, 2005.

[96]	H. Kassiri et al., "Rail-to-Rail-Input Dual-Radio 64-Channel Closed-Loop Neurostimulator," IEEE Journal of Solid-State Circuits, pp. 1-18, 2017, doi: 10.1109/jssc.2017.2749426.

[97]	M. Ghovanloo and K. Najafi, "A compact large voltage-compliance high output-impedance programmable current source for implantable microstimulators," IEEE Transactions on Biomedical Engineering, vol. 52, no. 1, pp. 97-105, 2004.

[98]	X. Liu, A. Demosthenous, and N. Donaldson, "An integrated implantable stimulator that is fail-safe without off-chip blocking-capacitors," IEEE Transactions on Biomedical Circuits and Systems, vol. 2, no. 3, pp. 231-244, 2008.

[99]	W.-M. Chen et al., "A fully integrated 8-channel closed-loop neural-prosthetic CMOS SoC for real-time epileptic seizure control," IEEE journal of solid-state circuits, vol. 49, no. 1, pp. 232-247, 2013.

[100]	J. F. Dickson, "On-chip high-voltage generation in MNOS integrated circuits using an improved voltage multiplier technique," IEEE Journal of solid-state circuits, vol. 11, no. 3, pp. 374-378, 1976.

[101]	A. Ballo, A. D. Grasso, and G. Palumbo, "A review of charge pump topologies for the power management of IoT nodes," Electronics, vol. 8, no. 5, p. 480, 2019.

[102]	M. Bugbee, N. de N Donaldson, A. Lickel, N. Rijkhoff, and J. Taylor, "An implant for chronic selective stimulation of nerves," Medical engineering & physics, vol. 23, no. 1, pp. 29-36, 2001.